Templates: Instrumental Descriptions for LC (Ultimate) QExactive HR-ESI-MS and -MS/MS

HR-ESI-MS with on-flow injection (routine service)

High-resolution electrospray ionization mass spectrometry (HR-ESI-MS): Dionex Ultimate 3000 UHPLC system (ThermoFischer Scientifics, Germering, Germany) connected to a QExactive MS with a heated ESI source (ThermoFisher Scientific, Bremen, Germany); on-flow injection of 1 L sample (c = ca. 50 g mL−1 in the indicated solvent) with an XRS auto-sampler (CTC, Zwingen, Switzerland); flow rate 120 L min−1; ESI: spray voltage 3.0 kV, capillary temperature 280 °C, sheath gas 30 L min−1, aux gas 8 L min−1, s-lens RF level 55.0, aux gas temperature 250 °C (N2); full scan MS in the alternating (+)/(−)-ESI mode; mass ranges 80–1’200 m/z, 133–2’000 m/z, or 200–3’000 m/z at 70’000 resolution (full width half-maximum); automatic gain control (AGC) target of 3.00·106; maximum allowed ion transfer time (IT) 30 ms; mass calibration <2 ppm accuracy with Pierce® ESI calibration solns. (ThermoFisher Scientific, Rockford, USA); lock masses: ubiquitous erucamide (m/z 338.34174, (+)-ESI) and palmitic acid (m/z 255.23295, (−)-ESI).

LC-HR-ESI-MS

Liquid chromatography high resolution electrospray ionization mass spectrometry (LC- HR-ESI-MS): Dionex Ultimate 3000 UHPLC system (ThermoFischer Scientifics, Germering, Germany) equipped with a DAD-3000 diode array detector and connected to a QExactive MS with a heated ESI source (ThermoFisher Scientific, Bremen, Germany); injection of 1 L sample (c = ca. 50 g mL−1 in the indicated solvent) with an XRS auto-sampler (CTC, Zwingen, Switzerland); Acquity BEH C18 HPLC column (1.7 m particle size, 2 × 100 mm, Waters) kept at 30 °C;* elution at a flow rate of 400 L min−1 with A: H2O + 1% HCO2H and B: CH3CN + 0.1% HCO2H, linear gradient from 5–98% B within 5 min, then isocratic for 1 min;* UV spectra recorded from 200–600 nm at 1.2 nm resolution and 20 points s−1; ESI: spray voltage 3.5 kV, capillary temperature 260 °C, sheath gas 45 L min−1, aux gas 15 L min−1, sweep gas 2 L min−1, s-lens RF level 450.0, aux gas temperature 250 °C (N2); full scan MS in (+)-ESI (or in (−)-ESI) mode; mass ranges 80–1’200 m/z, 133–2’000 m/z, or 200–3’000 m/z at 70’000 resolution (full width half-maximum)**; automatic gain control (AGC) target of 3.00·106; maximum allowed ion transfer time (IT) 30 ms; mass calibration <2 ppm accuracy with Pierce® ESI calibration solns. (ThermoFisher Scientific, Rockford, USA); lock masses: ubiquitous erucamide (m/z 338.34174, (+)-ESI) and palmitic acid (m/z 255.23295, (−)-ESI).

* check for effectively used chromatographic conditions (column and solvent system).

** check for effectively used resolution.